Title:

Simulated accuracy of transcutaneous CO₂ for hypercapnia detection using a meta-analytic error model

Authors:

Dustin Anderson-Bell MD

Jeanette Brown MD PhD

Ramkiran Gourripeddi MBBS MSc

? Ben Brintz

? Jason Ackrivo

Brian W Locke, MD MSCI

Background:

Hypercapnic respiratory failure is common and prognostically important, yet it is unreliably detected in clinical practice.1–6 The gold standard arterial blood gas (ABG) measurement of CO2 (PaCO2) requires arterial puncture expertise and can cause pain, vessel  injury, thrombosis, and blood stream infections. 7,8 This limits the role of ABGs for routine assessment for hypercapnia. Transcutaneous CO2 monitoring (TcCO2) is a potential non-invasive alternative.

Professional guidelines11 and prior work have suggested acceptable TcCO₂–PaCO₂ agreement thresholds of ±7.5 mm Hg and ±5.0 mm Hg, respectively, though these cutoffs reflect expert and clinician opinion rather than empirical evaluation of diagnostic consequences.  A systematic review and meta-analysis by Conway et al. estimated PaCO2-TcCO2 95% limits of agreements of -7.1 to 6.9 mmHg, but with uncertainty intervals encompassing disagreements as large as -15.1 to 14.9 mmHg, which has hampered adoption and reimbursement for TcCO2 assessment.9,10 However, the consequences of this magnitude of disagreement on diagnostic performance have not been evaluated. We therefore simulated hypercapnia misclassification rates using the Conway meta-analytic error model applied to the empirical distribution of PaCO₂ measurements from a large, nationwide electronic health record cohort.

Methods:

We queried the TriNetX research network (TriNetX LLC, Cambridge MA), which aggregates electronic medical records from approximately 115 million patients from 75 healthcare organizations). On 06/02/2023, we identified all adult encounters in 2022 with an arterial CO2 (PaCO2) measurement obtained on the first calendar day of care. Encounters were stratified by care location at the time of sampling: ambulatory (Amb), emergency department and inpatient (ED/Inp), and recipients of critical care (ICU).

To model agreement between TcCO₂ and PaCO₂, we utilized the systematic review and meta-analysis by Conway et al.10, which synthesized studies reporting paired TcCO₂–PaCO₂ measurements published prior to [DATE]. We implemented their mixed-effects Bland–Altman framework to estimate marginal (population) limits of agreement, defined as unconditional prediction intervals for TcCO₂–PaCO₂ differences that incorporate mean bias (μ), within-study variance (σ²), and between-study heterogeneity (τ²). These limits represent the expected disagreement for a new paired measurement in a new patient and clinical setting, rather than conditional agreement within a specific study or individual. We extended this model using bootstrap resampling to propagate uncertainty in μ, σ², and τ² into the resulting prediction intervals.

To estimate diagnostic misclassification, we simulated hypothetical TcCO₂ values for each encounter from TriNetX using the observed day-0 PaCO₂ as the reference, representing the TcCO₂ that would be expected had noninvasive monitoring been used instead of an arterial blood gas. Hypercapnia was defined as PaCO₂ ≥45 mm Hg, consistent with diagnostic criteria for hypercapnic respiratory failure and the recently proposed threshold for treatment response to domiciliary non-invasive ventilation in the US. Each simulated TcCO₂ measurement was classified as a true negative, false negative, false positive, or true positive based on concordance with the PaCO₂ threshold. We then calculated sensitivity, specificity, positive and negative likelihood ratios, and positive and negative predictive values, with bootstrap confidence intervals. These estimands describe the expected operating characteristics among patients who underwent arterial blood gas testing in participating institutions during 2022.

We additionally generated setting-specific conditional predictions by modeling care setting as a fixed effect: ambulatory, emergency department or inpatient ward, and intensive care. Ambulatory encounters were assumed to approximate the pulmonary function testing population in Conway et al., given the predominance of outpatient blood gases in that setting. Intensive care encounters were inferred from billing for critical care services in TriNetX, as explicit ICU location identifiers are not consistently available.

We also evaluated TcCO₂ thresholds of <40 mm Hg, 40–50 mm Hg, and >50 mm Hg, calculating interval likelihood ratios and classification rates to assess a two-stage strategy in which arterial blood gas testing is reserved for borderline TcCO₂ values.

Finally, because TcCO₂–PaCO₂ disagreement is symmetric, we derived prediction intervals for PaCO₂ conditional on an observed TcCO₂ value. We present the 95% prediction interval, the probability that PaCO₂ exceeded a specified threshold, and the most likely classification. We implemented two prediction approaches: (1) a likelihood-based method using only the meta-analytic error model from Conway et al., and (2) a Bayesian approach that combined this likelihood with the empirical PaCO₂ distribution from TriNetX to generate a posterior distribution of plausible PaCO₂ values given the observed TcCO₂. Conditional (setting-specific) predictions draw from the setting-specific observed PaCO2 values and treat setting as a fixed effect.

All analyses were performed in Python. Source code is publicly available at https://github.com/reblocke/tcco2-accuracy. An interactive web application for generating PaCO₂ prediction intervals from observed TcCO₂ values is available at https://tcco2-accuracy-2025-12-22.streamlit.app/. The codebase is structured to accommodate additional studies, enabling future updates consistent with a living systematic review. This study was reviewed and determined to be exempt from institutional review board oversight due to using only de-identified and publicly available data.

Results:

197,698 encounters with day-0 ABGs were included across all clinical settings (113,078 ED/Inpt, 66,704 ICU, and 10,456 Amb). Median age was 61 years (IQR 51, 74), 54.7% were male, and race was comprised of 64.5% non-Hispanic White, 14.3% Black, 5.4% Hispanic, and 1.9% Asian. Median PaCO2 for all aggregated patients was 40.0 (IQR 34.0, 46.6).  Prevalence of relevant comorbidities included 12.4% Obstructive Sleep Apnea (OSA), 11.5% asthma, 21.5% Chronic Obstructive Pulmonary Disorder (COPD), 22.4% Congestive Heart Failure (CHF), and 4.6% Opioid Use Disorder (OUD).

The proportion with PaCO2 ≥ 45 mmHg was 26.5% (Amb), 30.1% (ED/Inp), and 31.7% (ICU). [add Conway-derived agreement parameters we used: Mean bias (μ), Within-study variance (σ²), between-study heterogeneity (τ²), each with uncertainty intervals.]

Applying a definition of hypercarbia by TcCO2 of ≥45mmHg yielded the following operating characteristic estimates: Se: 81.6%-93.1%; Sp: 94.4%-96.9%; NPV: 92.4%-96.7%; and PPV: 85.2%-92.0%. The greatest false negative and false positive rates were 5.4% in the ED/Inp group and 3.8% in the Ambulatory group, respectively. With the more extreme TcCO2 cutoffs, values for Se/Sp/NPV/PPV all increased to >99% with false classification rates of <1% across all settings.

[Paragraph with “two-stage” results: the proportion of simulated TcCO₂ values in each zone (overall and by setting): % <40, % 40–50 (the “ABG reflex” zone), % >50. Then for each zone (and by setting), report: Interval-specific likelihood ratios (LR for TcCO₂ <40, 40–50, >50), Observed classification rates within each zone (e.g., % with PaCO₂ ≥45). And finally a concise operational interpretation: if ABG is reserved for 40–50, what fraction of patients would a.) require blood gas, b.) me misclassified? ]

[Paragraph for the prediction intervals: (example: 35, 40, 45, 50, 55 mm Hg), report: PaCO₂ 95% prediction interval. Posterior probability PaCO₂ ≥45 (or other relevant threshold) w side-by-side comparison of the two prediction approaches: Likelihood-only (meta-analysis model only) vs Bayesian (TriNetX prior + likelihood).]

Discussion:

Although TcCO₂ has not been shown to meet consensus agreement thresholds with high certainty, our simulation suggests that it would rarely misclassify the presence of hypercapnia among patients who currently undergo arterial blood gas testing. This is because only a small subset of patients have PaCO₂ values sufficiently close to the diagnostic threshold for plausible measurement disagreement to alter classification. Prior work has quantified population limits of agreement, which describe expected difference between paired TcCO₂ and PaCO₂ measurements; however, diagnostic misclassification depends on how that disagreement interacts with clinical thresholds. By applying an established meta-analytic error model to the real-world distribution of PaCO₂ values in a large, multi-institutional cohort, our study directly estimates hypercapnia misclassification rates if TcCO₂ were utilized.

Beyond diagnostic classification, these findings suggest several clinically relevant applications. First, TcCO₂ may support more patient-centered identification of de novo hypercapnia. Because arterial blood gas sampling is painful and invasive, clinicians typically reserve ABGs for situations with high pretest probability or anticipated clinical consequence, a practice that likely contributes to underrecognition. A reliable noninvasive alternative could enable earlier and more routine detection, analagous to pulse oximeters for hypoxemia detection.

Second, TcCO₂ may be useful for monitoring response to noninvasive ventilation, for which recent proposals recommend reassessment of hypercapnia at approximately six-month intervals. Although we could not model relative PaCO₂ reduction criteria (e.g., ≥20%) due to unknown baseline response rates, our results suggest TcCO₂ could serve as a practical longitudinal measure, consistent with current clinical use. Given low simulated misclassification rates, TcCO₂-based follow-up could avert many arterial blood gas measurements that may otherwise discourage ongoing engagement with chronic respiratory failure care. When higher accuracy is needed, a tiered strategy reserving ABG testing for borderline TcCO₂ values could further reduce invasive sampling while preserving diagnostic confidence.

This study has several strengths. We explicitly propagated multiple sources of uncertainty from a meta-analytic agreement model through to prediction intervals and diagnostic performance estimates, rather than relying on point estimates alone. We also grounded simulations in a large, multi-institutional cohort, reducing sensitivity to idiosyncratic blood gas ordering practices at individual sites. In addition, we provide an open, clinician-facing application that can be updated as new agreement data become available.

Important limitations remain. Our predictions condition on a limited set of factors, primarily care setting, whereas TcCO₂ accuracy is likely influenced by additional physiologic (e.g., perfusion), technical, and operator-level variables. The meta-analytic model from Conway et al. assumes normally distributed measurement error; however, large, non-normal discrepancies may occur, particularly if rates of operator or other pre-analytical error differ from those in the underlying studies. Binary identification of whether hypercapnia is present is only one use of TcCO2, while performance may differ for other uses like treatment response monitoring. Arterial blood gasses are obtained for other reasons than evaluating for hypercapnia so the spectrum of included patients may be overly broad. Broader use of TcCO₂ in patients with milder hypercapnia, such as for lower pre-test probability case finding or longitudinal treatment monitoring, would shift the PaCO₂ distribution toward diagnostic thresholds and likely increase misclassification. Most fundamentally, this work simulates classification error rather than directly measuring clinical performance, underscoring the need for future studies evaluating real-world deployment effects.

Conclusions

Despite wide population limits of agreement, our simulation suggests that TcCO₂ would rarely misclassify clinically significant hypercapnia when applied to patients who currently undergo arterial blood gas testing. These findings support a role for TcCO₂ as a noninvasive tool for detection hypercapnia and might be considered in clinical settings when arterial sampling is impractical or undesirable.

Funding

This work is supported by grants from the American Thoracic Society (ASPIRE fellowship), the Intermountain Foundation, and the National Institutes of Health (Ruth L. Kirschstein National Research Service Award 5T32HL105321 and National Center for Advancing Translational Sciences Award UM1TR004409).

References

1.	Konishi M, Akiyama E, Suzuki H, et al. Hypercapnia in patients with acute heart failure. ESC Heart Fail. 2015;2(1):12-19. doi:10.1002/ehf2.12023

2.	Nin N, Muriel A, Peñuelas O, et al. Severe hypercapnia and outcome of mechanically ventilated patients with moderate or severe acute respiratory distress syndrome. Intensive Care Med. 2017;43(2):200-208. doi:10.1007/s00134-016-4611-1

3.	Wilson MW, Labaki WW, Choi PJ. Mortality and Healthcare Use of Patients with Compensated Hypercapnia. Ann Am Thorac Soc. 18(12):2027-2032. doi:10.1513/AnnalsATS.202009-1197OC

4.	Fazekas AS, Aboulghaith M, Kriz RC, et al. Long-term outcomes after acute hypercapnic COPD exacerbation : First-ever episode of non-invasive ventilation. Wien Klin Wochenschr. 2018;130(19-20):561-568. doi:10.1007/s00508-018-1364-6

5.	Tiruvoipati R, Pilcher D, Buscher H, Botha J, Bailey M. Effects of Hypercapnia and Hypercapnic Acidosis on Hospital Mortality in Mechanically Ventilated Patients. Crit Care Med. 2017;45(7):e649-e656. doi:10.1097/CCM.0000000000002332

6.	Locke BW, Richards WW, Gouripeddi R, et al. The Consistency of Hypercapnic Respiratory Failure Case Definitions in Electronic Health Record Data. Chest. Published online August 28, 2025:S0012-3692(25)05126-8. doi:10.1016/j.chest.2025.08.002

7.	Baptistella CDPA, Mendes C de A, Silva MJ, Wolosker N. Retrospective Observational Single-Center Study of Complications of Arterial Indwelling Catheters for Invasive Blood Pressure in Intensive Care Unit Patients. Angiology. 2022;73(5):431-437. doi:10.1177/00033197211052127

8.	Scheer B, Perel A, Pfeiffer UJ. Clinical review: complications and risk factors of peripheral arterial catheters used for haemodynamic monitoring in anaesthesia and intensive care medicine. Crit Care Lond Engl. 2002;6(3):199-204. doi:10.1186/cc1489

9.	Farenden S, Chung Y, Cui A, Ferguson I. Prospective, observational study investigating the level of agreement between transcutaneous and invasive carbon dioxide measurements in critically ill emergency department patients. Emerg Med J EMJ. 2023;40(9):646-650. doi:10.1136/emermed-2022-212465

10.Conway A, Tipton E, Liu WH, et al. Accuracy and precision of transcutaneous carbon dioxide monitoring: a systematic review and meta-analysis. Thorax. 2019;74(2):157-163. doi:10.1136/thoraxjnl-2017-211466

11.Restrepo RD, Hirst KR, Wittnebel L, Wettstein R. AARC Clinical Practice Guideline: Transcutaneous Monitoring of Carbon Dioxide and Oxygen: 2012. Respir Care. 2012;57(11):1955-1962. doi:10.4187/respcare.02011